## Slide 1
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Large Data Sets
  Random variations will be present whenever we make a measurement:
      A large number of experiments done under identical conditions will yield a distribution of results.
      Equal chance of getting either high or low variations in a result – “bell-shape” curve centered around the average of the data set.
      Distribution of results is known as Normal distribution of a Gaussian distribution
Number of Occurrences
Value
High population about mean (m) or correct value
low population far from correct value
Width of ± 1 standard deviation (s)

## Slide 2
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Large Data Sets
  Normal Distribution:
      The shape of the Normal distribution of Gaussian Curve is described by the following equation:
        m is the average of the data set, which gives the central point for the distribution
        s is the standard deviation of the data set, which describes the width of this curve
      II. If our results follow a Normal distribution, we can use the average and standard deviation for the data set to determine what fraction of our results will fall between any two measured values.

## Slide 3
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
    The following table shows what fraction of results (as represented by the area under the Normal distribution will occur between the mean and a value x.

## Slide 4
Describing the Variation in Large Data Sets
  By knowing the standard deviation (s) and the average (m):
      The probability of the next result falling in any given range can be calculated by:
        Describes the difference between x and m in terms of the number of standard deviations that separate these two values.
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Standard Deviation (s) | Probability
±1s | 68.3%
±2s | 95.5%
±3s | 99.7%
±4s | 99.9%
Probability of Measuring a value in a certain range is equal to the area of that range
As an example, a range of one standard deviation above or below the mean (m±1s) corresponds to a relative area of 2(0.3413) = 0.6826 or 68.3% of the results in a normal distribution, or roughly two thirds of all its values

## Slide 5
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Small Data Sets
  For a small set of numbers:
      The experimental values of and s are only estimates of the true average (m) and standard deviation (s).
      We must always consider how precisely we know and s when we use these to describe experimental data.
  Standard Deviation of the Mean:
      In the same way that we use s to describe the variation within a data set, we can employ the standard deviation of the mean () to describe the precision of our experimental average ()
      The standard deviation of the mean () is determined by using the standard deviation of the entire data set (s) and the number of data points (n) in this data set:

## Slide 6
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Small Data Sets
  Standard Deviation of the Mean:
      The size of is always less than or equal to s since n must be greater than or equal to one
      Whenever a standard deviation for a mean is reported, you also need to state the number of points in your data set - depends on n
Chart: 
 | 1 | 0.70711 | 0.57735 | 0.5 | 0.44721 | 0.40825 | 0.37796 | 0.35355 | 0.33333 | 0.31623 | 0.30151 | 0.28868 | 0.27735 | 0.26726 | 0.2582 | 0.25 | 0.24254 | 0.2357 | 0.22942 | 0.22361
Chart: 
 | 1 | 0.70711 | 0.57735 | 0.5 | 0.44721 | 0.40825 | 0.37796 | 0.35355 | 0.33333 | 0.31623 | 0.30151 | 0.28868 | 0.27735 | 0.26726 | 0.2582 | 0.25 | 0.24254 | 0.2357 | 0.22942 | 0.22361
Number of Assays Required (n)
Relative size of versus s

## Slide 7
Describing the Variation in Small Data Sets
  Standard Deviation of the Mean:
      The standard deviation for an entire set (s) approaches a constant value (s) as we increase n
      The size of becomes smaller as we increase n
        This occurs because the precision of the experimental average decreases as we acquire more data
        x is a more reliable estimate of the true average as n increases
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Chart: 
 | 1 | 0.70711 | 0.57735 | 0.5 | 0.44721 | 0.40825 | 0.37796 | 0.35355 | 0.33333 | 0.31623 | 0.30151 | 0.28868 | 0.27735 | 0.26726 | 0.2582 | 0.25 | 0.24254 | 0.2357 | 0.22942 | 0.22361
Chart: 
 | 1 | 0.70711 | 0.57735 | 0.5 | 0.44721 | 0.40825 | 0.37796 | 0.35355 | 0.33333 | 0.31623 | 0.30151 | 0.28868 | 0.27735 | 0.26726 | 0.2582 | 0.25 | 0.24254 | 0.2357 | 0.22942 | 0.22361
Number of Assays Required (n)
Relative size of versus s

## Slide 8
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Small Data Sets
  Confidence Intervals:
      Common in science to describe the variation in experimental numbers by using a range of values
        Report a result by giving the mean plus or minus two standard deviations of the mean:
        The range of values that follows the mean is called the confidence limit
        The mean plus this range is known as the confidence interval (or C.I.)

## Slide 9
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Small Data Sets
  Confidence Intervals:
      II. When reporting a confidence interval, the number placed in front of helps specify the degree of certainty that the experimenter has in the result
        For a Normal distribution, a range of approximately ± 2 standard deviations means there is roughly a 95% chance that any given value in the data set will fall in this range
        Only a 5% chance that a value will fall outside of this range
-2s +2s
95% of the area of a Normal distribution or Gaussian curve

## Slide 10
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Describing the Variation in Small Data Sets
  Confidence Intervals:
      III. Relatively easy to determine the meaning of these ranges for large groups of numbers, this becomes more complicated for small data sets
        Mean and standard deviations are only estimates of their true value
        Always a greater uncertainty when working with small data sets
        Requires the use of larger confidence intervals
      Use a correction factor known as the Student’s t value (t)
        Express the confidence interval for an entire population of results based on s
        Express the confidence interval for the measurement of a mean based on

## Slide 11
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
    The following table gives the Student’s t values for a given number of points (n) in your data set as represented by the degrees of freedom (n ).
Student’s t value also depends on the desired degree of certainty – Confidence level
As n becomes large, t approaches ~ 2
(± 2 standard deviations)

## Slide 12
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Calculating a Confidence Interval
      Probenecid is a drug used by some athletes to prevent the excretion of other substances into urine, thus lowering their detectable concentrations. A scientist makes three measurements of a urine sample known to contain probenecid. He gets a mean result of 11.8 mg/L and a standard deviation for the entire set of results of 0.2 mg/L
      What is the 95% confidence interval for this mean?
Solution: Since we are looking at the mean, we first need to find
Next, look-up the Student’s t value from the table at the correct degrees of freedom of 2
at the 95% confidence level, t = 4.303

## Slide 13
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Calculating a Confidence Interval
      Probenecid is a drug used by some athletes to prevent the excretion of other substances into urine, thus lowering their detectable concentrations. A scientist makes three measurements of a urine sample known to contain probenecid. He gets a mean result of 11.8 mg/L and a standard deviation for the entire set of results of 0.2 mg/L
      What is the 95% confidence interval for this mean?
Solution: With = 0.12 mg/L and t = 4.303, we can now calculate the confidence interval:
Note: always state the number of data points and confidence level when a confidence
interval is reported

## Slide 14
Comparing Experimental Results
  General Requirements for the Comparison of Data:
  There are four items you need when using statistics to compare experimental results:
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
The TEST STATISTIC
“How will I compare my result and model?”
The CONFIDENCE LEVEL
“How certain do I want my answer to be?”
The HYPOTHESIS
“Is my result the same as the model?”
The MODEL
“What is my result being compared to?”

## Slide 15
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  The MODEL - refers to the value or predicted behavior to which the experimental results are going to be compared.
    This could be an equation, a predicted distribution, the values obtained by another method, or the known value for a reference standard.
  The HYPOTHESIS - is an initial guess for the results of the statistical test..
    When comparing analytical results, the hypothesis can be either:
      the results will fit the model – the null hypothesis
      the results will not fit the model – the alternate hypothesis
  The CONFIDENCE LEVEL - represents the degree of certainty required in the comparison.
    Scientific results have some degree of uncertainty because of random errors
    Confidence level estimates the extent of this uncertainty to avoid reaching unreasonable conclusions about the data.
  The TEST STATISTIC - a numerical value calculated from the data to use in the comparison (e.g., Student’s t value).
    The test statistic calculated from the results is compared to a critical value that represents the largest value is expected for a given number of data points and confidence level.

## Slide 16
Comparing Experimental Results
  Comparing an Experimental Result with a Known Reference Value:
  If the reference value is known exactly, or at least has much better precision than the results, then
    The known reference value represents the true “mean” for the sample, m
    The experimental results is the measured mean for the sample, x
    The Student’s t value is the test statistic
Student’s t test
  Assume the reference value (m) and the experimental results (x) are the same – the null hypothesis
  Test this assumption by calculating a Student’s t value
  The lines shown on either side of “x - m” indicate the absolute or positive value of the difference
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 17
Comparing Experimental Results
  Comparing an Experimental Result with a Known Reference Value:
  III. Once t is calculated for the data, need to compare this to a critical value (tc) obtained from a table of expected Student’s t values
    The selected tc value is determined by the number of data points (n) used to find the experimental mean (degrees of freedom, n-1)
    The selected tc value is determined by the confidence level chosen for the comparison
IV. If t ≤ tc, then x and m are not significantly different at the stated confidence level.
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 18
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Comparing an Experimental Result and a Known Reference Value
      Action is taken against Olympic athletes if their urine is found to contain caffeine concentrations above 12.00 mg/mL. A sample from one athlete gives a mean caffeine concentration of 11.85 mg/mL for five measurements (range, 11.65 to 12.10 mg/mL) with a standard deviation for this mean being 0.07 mg/mL. The athlete's coach argues that this result is statistically the same as the 12.00 mg/mL cutoff.
      Are these two values equivalent at the 95% confidence level?
Solution: The model in this example is 12.00 mg/mL, and the confidence level is 95%. To see if the mean and the reference value are the same (underlying hypothesis) calculate the Student’s t value:
Next, look-up the critical Student’s t value from the table at the 95% confidence level and at a degrees of freedom of (5-1)=4

## Slide 19
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Comparing an Experimental Result and a Known Reference Value
      Action is taken against Olympic athletes if their urine is found to contain caffeine concentrations above 12.00 mg/mL. A sample from one athlete gives a mean caffeine concentration of 11.85 mg/mL for five measurements (range, 11.65 to 12.10 mg/mL) with a standard deviation for this mean being 0.07 mg/mL. The athlete's coach argues that this result is statistically the same as the 12.00 mg/mL cutoff.
      Are these two values equivalent at the 95% confidence level?
Solution: at the 95% confidence level and v = 4, tc = 2.776
Since tc of 2.776 is greater than the experimental t value of 2.14, the amount of caffeine in the athlete’s sample was the same as the allowed cutoff level at a 95% confidence level

## Slide 20
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  Comparing Two Experimental Means:
  Mean results for two samples (x1 and x2) measured by the same method or two methods with similar precision – are they the same?
    The model would be one of the two means
    The hypothesis is determining if the two results represent the same number
    The Student’s t value is the test statistic
Pooled Standard Deviation (spool)
  Both experimental result and “model” have some uncertainty in their values
  Instead of using standard deviation for either of these means, the pooled standard deviation (spool) reflects the variation in both results
  where
  s1 and s2 are the estimated standard deviations for the two datasets
  n1 and n2 are the number of points for the two datasets

## Slide 21
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  Comparing Two Experimental Means:
Pooled Standard Deviation (spool)
  spool is the weighted average of the individual standard deviations
  Just like s can be used to find the standard deviation of the mean x, spool can be used to determine the standard deviation for the pooled mean ( ):
  If x1 and x2 represent the same value, their difference (x1 - x2 ) should fall within a reasonably small number of standard deviations for this difference
    (x1 - x2 ) can be compared directly to , where a ratio gives an experimental Student’s t value for the dataset:

## Slide 22
Comparing Experimental Results
  Comparing Two Experimental Means:
  IV. Once t is calculated for the data, need to compare this to a critical value (tc) obtained from a table of expected Student’s t values
    The selected tc value is determined by the number of data points (n1 and n2) used to find the experimental means (degrees of freedom, n1+n2-2)
    The selected tc value is determined by the confidence level chosen for the comparison
V. If t ≤ tc, then x1 and x2 represent the same value at the stated confidence level.
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 23
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Comparing Two Mean Results
      Human chronic gonadotropin (hCG) is a naturally-occurring substance that has been abused by some athletes because of its ability to stimulate testosterone production. Two labs that perform athletic drug testing are to be evaluated for their ability to measure this hormone by using the same sample and analysis method. The first lab reports a mean hCG level of 2.99 IU/L (n1 = 4) with a standard deviation of 0.06 IU/L, while the second lab obtains a mean level of 3.13 IU/L (n2 =5) with a standard deviation of 0.08 IU/L
      Are these mean results the same at the 95% confidence level?
Solution: If we assume that the standard deviations for the two means are approximately the same, the first step is to get the pooled standard deviation:

## Slide 24
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution: Next, we can use spool , n1 and n2 to determine the standard deviation of the pooled mean ():
We are now ready to calculate the experimental Student’s t value for our results:

## Slide 25
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution: The degrees of freedom in this case is:
(4 + 5 -2) = 7
At a 95% confidence level, the critical tc value is 2.36
When we compare the experimental t value and the critical tc value:
t is greater than tc (2.9 > 2.36)
The mean results from the two labs are significantly different at the 95% confidence level

## Slide 26
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  Comparing Two Sets of Experimental Data:
  Mean results for one sample measured by two different methods (x1 and x2) – are they the same?
    The two methods need to have the same precision
Paired Student’s t test
  Make a list of the results obtained by both methods for each sample
 | Mean results (mmol/L) |  | Difference in Results (mmol/L)
Sample No. | Method 1 (x1) | Method 2 (x2) | di = x1-x2
1 | 2.53 | 2.68 | -0.15
2 | 5.19 | 5.03 | 0.16
3 | 3.60 | 3.79 | -0.19
4 | 6.42 | 6.51 | -0.09
5 | 7.08 | 7.24 | -0.16
 |  |  | = -0.086 mmol/L
 | Mean results (mmol/L) |  | Difference in Results (mmol/L)
Sample No. | Method 1 (x1) | Method 2 (x2) | di = x1-x2
1 | 2.53 | 2.68 | -0.15
2 | 5.19 | 5.03 | 0.16
3 | 3.60 | 3.79 | -0.19
4 | 6.42 | 6.51 | -0.09
5 | 7.08 | 7.24 | -0.16

## Slide 27
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  Comparing Two Sets of Experimental Data:
Paired Student’s t test
  The difference between each set of results is calculated (di)
  The average of the differences in the results is averaged ()
  To determine whether the differences in these results are significant, we need to calculate the standard deviation in these differences (sd):
  Next, calculate the standard deviation in the average difference ()

## Slide 28
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  Comparing Two Sets of Experimental Data:
Paired Student’s t test
  If the differences in the results for methods one and two represent only random variations, then the average difference in these results should be similar in size to .
  The experimental Student’s t value for this analysis is calculated as follows:
  Compare the experimental t value, to a critical Student’s tc value
    Required confidence level at n - 1 degrees of freedom
    N now represents the number of data point pairs being compared
  If t ≤ tc, the two methods produce statistically identical values at the given confidence level

## Slide 29
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Paired Student’s t Test
      Corticosteroids can legitimately be used by athletes for the relief of inflammation and pain. But the injection or inhalation of these compounds is allowed only when needed for a medical condition. A new technique for the measurement of corticosteroids in urine is to be compared with a previous method. Both approaches have similar precision and are used to analyze a series of identical samples. The new method gives mean results of 2.53, 5.19, 3.60, 6.42, and 7.08 mmol/L for five separate samples, while the older method gives means of 2.68, 5.03, 3.79, 6.51 and 7.24 mmol/L for the same samples.
      Are the results from these methods equivalent at the 95% confidence level?
Solution: Use a paired Student’s t test, and first list the results for all samples side-by-side:
 | Mean results (mmol/L) |  | Difference in Results (mmol/L)
Sample No. | Method 1 (x1) | Method 2 (x2) | di = x1-x2
1 | 2.53 | 2.68 | -0.15
2 | 5.19 | 5.03 | 0.16
3 | 3.60 | 3.79 | -0.19
4 | 6.42 | 6.51 | -0.09
5 | 7.08 | 7.24 | -0.16
 |  |  | = -0.086 mmol/L
 | Mean results (mmol/L) |  | Difference in Results (mmol/L)
Sample No. | Method 1 (x1) | Method 2 (x2) | di = x1-x2
1 | 2.53 | 2.68 | -0.15
2 | 5.19 | 5.03 | 0.16
3 | 3.60 | 3.79 | -0.19
4 | 6.42 | 6.51 | -0.09
5 | 7.08 | 7.24 | -0.16

## Slide 30
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution: Calculate the difference between each pair of results, and then the average difference between the two methods: -0.086 mmol/L
Calculate sd:
Calculate :

## Slide 31
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution: Calculate the experimental Student’s t value
The degrees of freedom in this case is: n - 1 = 5 - 1= 4
At a 95% confidence level, the critical tc value is 2.78
When we compare the experimental t value and the critical tc value:
t is less than tc (1.4 < 2.78)
The results from the two methods are equivalent at the 95% confidence level

## Slide 32
Comparing Experimental Results
  Comparing the Variation in Results:
  Compare the precision of two results or methods
    The methods we have discussed to this point require similar precision
F test
  The model is the method or result with the smallest standard deviation (s1)
  The hypothesis is that the standard deviation from the second method or result (s2) is the same as the model’s standard deviation (s1)
  The test statistic is the ratio of the squared standard deviations:
  Since s1 < s2, F should always be greater than or equal to one
  As F becomes larger, there is a greater likelihood that s1 and s2 represent different numbers
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 33
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Comparing Experimental Results
  Comparing the Variation in Results:
F test
  After F is calculated for the data set, it needs to be compared to an appropriate critical value - Fc
  The Fc value is determined by the desired confidence level
  The Fc value is determined by the degrees of freedom:
      v1 = n1 -1 and v2 = n2 -1
  where n1 and n1 are the number of points for data sets one and two
  If F ≤ Fc, the precision of the two methods is equivalent at the selected confidence level

## Slide 34
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 35
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Comparing the Precision of Two Methods by the F-Test
      It is known that the two methods in the previous example have standard deviations of 0.09 and 0.16 mmol/L (for n1 = n2 = 5) at a corticosteroid concentration of 5.0 mmol/L
      Are the precisions of these two methods the same at the 95% confidence level?
Solution: Set s2 equal to 0.16 and s1 equal to 0.09, so that s2 > s1, and determine F:
From the Table on the previous slide, the Fc critical value at the 95% confidence level for the two degrees of freedom (v1 = n1-1 = 5-1 =4 and v2 = n2-1 = 5-1 =4) is 6.39
Since F ≤ Fc (3.2 ≤ 6.39), the two methods have the same precision at the 95% confidence level.

## Slide 36
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Detecting Outliers
  Variations are Always Present for Repeated Measurements on a Sample:
  A data point that is very different from others obtained under supposedly identical conditions is suspect
    Is this due to a problem with the experiment?
    Experience can be used to identify an obviously erroneous data point and remove it from the dataset
    There are other occasions when experience is not sufficient, the data point doesn’t appear to fit the general trend for other results, but is it an outlier?
Various tests for determining if a data point is outside the variation normally expected for a dataset
  Only used for identifying outliers
  Not the sole means for justifying the removal of a data point
  Thorough knowledge of the methods and conditions should always have the “last word” in determining whether a point should be kept in a data set

## Slide 37
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Detecting Outliers
  Variations are Always Present for Repeated Measurements on a Sample:
  III. Q Test
    Based on the absolute difference between a suspect data point’s value and the nearest data point. This difference is then compared to the total range of values in the data set.
    If the difference between the suspect data point and its nearest neighbor is greater than a certain critical fraction of the total range, then the suspected value is a “true” outlier
IV. Application of the Q test
  Rank the results from the data set from lowest to highest
  Define the suspected outlier xo and its nearest neighbor xn
  Define the highest number (xhigh) and the lowest number (xlow) in the data set
  Calculate the following ratio (Q):

## Slide 38
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Detecting Outliers
  Variations are Always Present for Repeated Measurements on a Sample:
  IV. Application of the Q test
  Compare the calculated value for Q to a critical test value, Qc
  The critical test value will depend on the total number of results in the data set
  The critical test value will depend on defined confidence level
  If Q > Qc, the suspected point can be called an outlier and considered for rejection

## Slide 39
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
 | Values for Qc at Various Confidence Levels |  | 
Number of Values in Data Set | 90% | 95% | 99%
3 | 0.941 | 0.970 | 0.994
4 | 0.765 | 0.829 | 0.926
5 | 0.642 | 0.710 | 0.821
6 | 0.560 | 0.625 | 0.740
7 | 0.507 | 0.568 | 0.680
8 | 0.468 | 0.526 | 0.634
9 | 0.437 | 0.493 | 0.598
10 | 0.412 | 0.466 | 0.568
11 | 0.392 | 0.444 | 0.542
12 | 0.376 | 0.426 | 0.522
13 | 0.361 | 0.410 | 0.503
14 | 0.349 | 0.396 | 0.488
15 | 0.338 | 0.384 | 0.475
16 | 0.329 | 0.374 | 0.463
17 | 0.320 | 0.365 | 0.452
18 | 0.313 | 0.356 | 0.442
19 | 0.306 | 0.349 | 0.433
20 | 0.300 | 0.342 | 0.425

## Slide 40
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Outlier Detection by the Q Test
      A urine sample containing a known amount of markers for marijuana is sent to several drug testing labs to evaluate their ability to monitor such compounds. These labs report the following concentrations: Lab 1: 55.3 mg/L, Lab 2: 57.8 mg/L, Lab 3: 54.0 mg/L, lab 4: 68.1 mg/L, and Lab 5 58.7 mg/L
      Use the Q test to determine if any of these results can be considered an outlier at the 95% confidence level
Solution: The low and high values in the group are 54.0 and 68.1 mg/L. The result of 68.1 mg/L is the most likely outlier since it is the furthest from its neighbor, 58.7 mg/L.
Calculate a Q value:

## Slide 41
      Example – Outlier Detection by the Q Test
      A urine sample containing a known amount of markers for marijuana is sent to several drug testing labs to evaluate their ability to monitor such compounds. These labs report the following concentrations: Lab 1: 55.3 mg/L, Lab 2: 57.8 mg/L, Lab 3: 54.0 mg/L, lab 4: 68.1 mg/L, and Lab 5 58.7 mg/L
      Use the Q test to determine if any of these results can be considered an outlier at the 95% confidence level
Solution: The number of points in this data set is 5 – note: it is not a degrees of freedom
At a 95% confidence level, the critical Qc value is 0.710 from the table on the previous slide
Since the calculated Q is less than the critical Qc value ( (0.667 < 0.710), the point at 68.1 mg/L can not be called an outlier at the 95% confidence level
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 42
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Fitting Experimental Results
  Linear Regression:
  How to fit an equation or line to a set of results
    Many types of equations, but the most common is a straight line
    A common method for deterring the best-fit line for a data set is a process known as linear regression
Application of Linear Regression
  Involves a set of (x, y) values
  y is the dependent variable, and x is the independent variable
  Fit to an equation with the following form:
    where:
      m is the slope (representing the change in y versus x)
      b is the line’s intercept on the y-axis
      xi is a given x value in the data set
      yi,calc is the response predicted at xi by the best-fit line

## Slide 43
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Fitting Experimental Results
Application of Linear Regression
  Obtain the best estimates for m and b by using the method of least squares analysis.
  Least squares analysis results in a series of equations that allow the slope and intercept for the best-fit line to be calculated for a particular data set based on the number of points in the data set (n) and the values for each (x, y) pair
  Can be calculated manually, best-fit lines are routinely determined using a computer

## Slide 44
Fitting Experimental Results
  III. Least Squares Analysis.
    Minimize vertical deviation between points and line
    Use square of the deviations  deviation irrespective of sign
STATISTICAL TESTS FOR CHEMICAL ANALYSIS

## Slide 45
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Determining the Best-Fit Parameters for a Line
      A set of urine standards that contain the drug oxymorphone are analyzed and give a calibration curve that appears to follow a straight line. The peak heights measured by liquid chromatography for standards with oxymorphone concentrations of 100, 200, 300, 400, and 500 ng/mL have relative values of 161, 342, 543, 765, and 899, respectively.
      Determine the best-fit slope and intercept for this line.
Solution: The easiest approach to solve this problem is to prepare a table, which has separate columns for each x and y pair, as well as for the calculated values of and . The numbers in each column are then summed:

## Slide 46
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution: Table of data, calculated values, and sums:
The best-fit slope (m), can now be calculated from these sums:
Drug Conc. (x) | Peak Height (y) | xiyi | xi2
100 | 161 | 16,100 | 10,000
200 | 342 | 68,400 | 40,000
300 | 543 | 162,900 | 90,000
400 | 765 | 306,000 | 160,000
500 | 899 | 449,500 | 250,000
1500 | 2710 | 1,002,900 | = 550,000
Drug Conc. (x) | Peak Height (y) | xiyi | xi2
100 | 161 | 16,100 | 10,000
200 | 342 | 68,400 | 40,000
300 | 543 | 162,900 | 90,000
400 | 765 | 306,000 | 160,000
500 | 899 | 449,500 | 250,000

## Slide 47
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution:
Similarly, we can use these sums to get the best-fit intercept (b):
Drug Conc. (x) | Peak Height (y) | xiyi | xi2
100 | 161 | 16,100 | 10,000
200 | 342 | 68,400 | 40,000
300 | 543 | 162,900 | 90,000
400 | 765 | 306,000 | 160,000
500 | 899 | 449,500 | 250,000
1500 | 2710 | 1,002,900 | = 550,000
Drug Conc. (x) | Peak Height (y) | xiyi | xi2
100 | 161 | 16,100 | 10,000
200 | 342 | 68,400 | 40,000
300 | 543 | 162,900 | 90,000
400 | 765 | 306,000 | 160,000
500 | 899 | 449,500 | 250,000

## Slide 48
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Solution: Thus, the best-fit line to the data set is y = 1.90x - 28
Chart: Drug Concentration
| 100 | 200 | 300 | 400 | 500
 | 161 | 342 | 543 | 765 | 899
Results from
Microsoft Excel

## Slide 49
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Fitting Experimental Results
  IV. Formulas for Determining the Best-Fit Parameters for a Straight Line
Equation for a Line:
Slope (m):
Intercept (b):
Standard deviation of y values (sy):
Standard deviation of slope(sm):

## Slide 50
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Fitting Experimental Results
  IV. Formulas for Determining the Best-Fit Parameters for a Straight Line
Standard deviation of intercept (sb):
Correlation coefficient (r):
where:

## Slide 51
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Fitting Experimental Results
  V. Testing the Goodness of a Fit.
    Given a best-fit line, it is essential to check and make sure that it does present a good description of the data – known as “goodness of fit”
    Correlation coefficients and residual plots are used to determine the goodness of fit
    VI. Correlation coefficient (r)
    Indicates how well a best-fit line describes the data
    Equations on previous slide
    Gives a value between -1 and 1
      coefficient of determination (r2) is the square of the correlation coefficient and gives a value between 0 and 1
    A value of r equal to 1 or -1 represents a perfect agreement between the data points and the best-fit line
    A value of r equal to 0 represents a random relationship between the data points and the best-fit line
    A positive value for r means y and x are changing in the same direction
    A negative value for r means y and x are changing in opposite direction

## Slide 52
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
      Example – Determining the Correlation Coefficient for a Best-Fit Line
      What is the correlation coefficient for the best-fit line to the calibration curve in the previous example? What is the probability that this line represents a real trend between the x and y values in this data set?
Solution: The correlation coefficient for this data can be calculated using the prior equations. This, in turn, requires that we first use the equations to find sxy, sxx, and syy. The values of in these equations are obtained from the table previously created. In the same way, can be calculated from the table, given a value of 1,831,160. These values can then be used to determine sxy, sxx, and syy.

## Slide 53
Solution:
These values are then used to calculate the correlation coefficient (r):
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
The variation in y accounts for 99.8% of the variation in x

## Slide 54
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Fitting Experimental Results
    VII. Residual Plot
    Although the correlation coefficient gives some indication as to how well a line fits a set of data, this should not be used alone in determining the goodness of fit
    There are many cases where a good correlation coefficient is obtained, but the data does not really fit the line
      Residual plot detects and avoids this problem
Application of the Residual Plot
    Plot the difference or residual between each experimental value for the dependent variable (yi) and the value predicted by the best-fit line (yicalc).
    Include a reference line that shows where (yi - yicalc) = 0, the result for a perfect agreement between the data and best-fit line
    If the best-fit line is a good description of the data, the residual plot should only have a random distribution of points above and below the line at (yi - yicalc) = 0
    If the best-fit line is a poor description of the data, then a definite trend in the residual points should appear – an alternative fit is needed.
    Can be used with other best-fit equations besides the equation for a straight-line.

## Slide 55
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Original Plots
Residual Plots
Method 1
Method 2
Good Fit
Poor Fit

## Slide 56
STATISTICAL TESTS FOR CHEMICAL ANALYSIS
Learning Objectives
    Be able to describe what is meant by a normal distribution, the general factors used to describe such a distribution, and methods for determining the probability that a given result will occur in a particular range of such a distribution.
    Be able to define and calculate/use each of the following terms:
    Standard deviation of the mean Confidence interval
    Student’s t value Confidence level
    3. Be able to describe the four items needed when using statistics to compare experimental results.
    4. Be familiar with each of the following statistical tests and their use in comparing or evaluating experimental results:
    Student’s t test Paired Student’s t test
    F test Q test
    5. Be able to discuss the process of linear regression and be able to perform the necessary calculations when using this method for a set of data.
    6. Be able to use correlation coefficients and residual plots for testing the goodness of the fit of a line to data